## Slide 1
Presentation :
Specifications:
    19 inch half-cabinet
    Height : 2U
    Width: 215 mm
    Depth: 265 mm
Minimum volume and weight for maximum capacity
Switched mode technology
Voltage and current controlled by PWM (Pulse Width Modulation)
On request : programmable voltage/current: 0-10 V, isolated to 1500 Vrms
Over-voltage protection (OVP) / Over-temperature protection (OTP)
U
I
P
OVP
OT
USB
RS232
LAN
IEEE
CAN
Profibus
DC OUTPUT (max 300W) : 0 - 20 Amps
0 - 1200 Volts
Stabilized laboratory power-supply
MICROLAB DC MX 300W
Micronics Systems l Z.A. 18 Rue des Emeraudes l T +33 (0) 472 930 480 l contact@micronics.fr
France l F-38280 Villette d’Anthon l F +33 (0) 472 930 481 l www.micronics.fr
Power : MX 300W Weight : 3 Kg | 
Volts | Amps
0 - 15 | 0 - 20
0 - 30 | 0 - 10
0 - 50 | 0 - 6
0 - 75 | 0 - 4
0 - 100 | 0 - 3
0 - 150 | 0 - 2
0 - 180 | 0 – 1,6
0 - 250 | 0 – 1,2
0 - 350 | 0 – 0,8
0 - 750 | 0 – 0,4
0 - 1200 | 0 - 0,25

## Slide 2
Customisation :
Digital control:
  2-wire RS 485 / Ethernet /
  Wifi
  Profibus DP
  Standard : Schneider/
  Siemens/Wago/Phoenix…
Analog control:
  U and I Inputs/Outputs
  0-10V or 4-20 mA isolated 1500 Vrms
15-pin D-Sub connector for remote programming
Descriptions :
. Polycarbonate front panel
. Two-colour On/Off LED: Off = red / On = green
. Voltage and current adjustable using two 10-turn potentiometers (U-adj and I-adj)
. Two modes of operation: constant current (CC) and constant voltage (CV) indicated by two LEDs
. 2000-point digital displays (VOLTAGE/CURRENT)
. Power output activated by pressing 'OUT'
. Manual Voltage/Current preset (U/I set push button) shown on displays
. Power input by standardised pluggable connector
. Power outputs on insulated terminals or screw terminal block
Non-binding document, may be amended without notice.
Ref: 140217 / MX 300W
Stabilized laboratory power-supply
MICROLAB DC MX 300W
Micronics Systems l Z.A. 18 Rue des Emeraudes l T +33 (0) 472 930 480 l contact@micronics.fr
France l F-38280 Villette d’Anthon l F +33 (0) 472 930 481 l www.micronics.fr
Power : MX 300W Weight : 3 Kg |  | 
AC Input |  | 
Voltage | 230 Vac Mono + 10% -15% | 
Frequency | 47 / 63 Hz | 
Power factor | > 0,95 | 
DC Output |  | 
Power |  | 300 W
Voltage |  | 0 to 1200 V
Current |  | 0 to 20 A
Mode of operation |  | CV/CC
Residual ripple |  | < 0,5% of the normal gauge
Accuracy ( voltage and current ) |  | < 0,2% of the normal gauge
Stability |  | < 0,2% of the normal gauge
Linearity |  | < 0,2% of the normal gauge
Distortion |  | < 0,5%
Environment |  | 
Temperature | 0 to +40°C ( 55°C derating ) | 
Relative Humidity | 85% without condensation | 
Cooling | Controlled air cooling | 
Dielectric Rigidity | 2500 Vrms | 
ON / OFF
MANUAL SETTING
“OUT” BUTTON
1
2
3
4
“PRG” BUTTON
1
2
3
4
DC OUTPUT
COOLING FAN
D-SUB CONNECTOR
Remote analog control
5
6
7
8
POWER INPUT
Controlled air cooling
0 to 1200V / 0 to 20A
230 Vac Mono
5
6
7
8